## Slide 1: BMS 631: Lecture 4
Fluorescence
J. Paul Robinson, PhD
The SVM Professor of Cytomics &
Professor of Biomedical Engineering
Purdue University
www.cyto.purdue.edu
All materials used in this course are available for download on the web at http://tinyurl.com/2wkpp
Last updated Feb, 2022
Notice: The materials in this presentation are copyrighted materials. If you want to use any of these slides, you may do so if you credit each slide with the author’s name. It is illegal to place these notes on CourseHero.

## Slide 2: Absorption
Basic quantum mechanics requires that molecules absorb energy as quanta (photons) based upon a criteria specific for each molecular structure
Absorption of a photon raises the molecule from ground state to an excited state
Total energy is the sum of all components (electronic, vibrational, rotational, translations, spin orientation energies) (vibrational energies are quite small)
3rd Ed Shapiro p 84
Absorption spectrum of FITC

## Slide 3: Fluorophores have different absorption properties
The structure of the molecule dictates the likely-hood of absorption of energy to raise the energy state to an excited one
Calcein
FITC

## Slide 4: Lifetime
Absorption associated with electronic transitions (electrons changing states) occurs in about 1 femtosecond (10-15 s)
The lifetime of a molecule depends on how the molecule disposes of the extra energy – but many fluorophores have lifetimes around 5 ns
Because of the uncertainty principle, the more rapidly the energy is changing, the less precisely we can define the energy
So, long-lifetime-excited-states have narrow absorption peaks, and short-lifetime-excited-states have broad absorption peaks
3rd Ed. Shapiro p 85
4th Ed. Shapiro p 110
Notes: Its useful to understand that the lifetime of many of the fluorophores is very short – when I say short, I mean really short – many are around 5-10 nanoseconds…

## Slide 5: Fluorescence Lifetime
Absorption associated with electronic transitions (electrons changing states) occurs in about 1 femptosecond (10-15 s)
Fluorescence lifetime is defined as the time in which the initial fluorescence intensity of a fluorophore decays to 1/e (approx 37 percent) of the initial intensity
The lifetime of a molecule depends on how the molecule disposes of the extra energy
Because of the uncertainty principle, the more rapidly the energy is changing, the less precisely we can define the energy
So, long-lifetime-excited-states have narrow absorption peaks, and short-lifetime-excited-states have broad absorption peaks
Shapiro p 85
http://micro.magnet.fsu.edu/primer/techniques/fluorescence/fluorescenceintro.html

## Slide 6: Fluorescence Lifetimes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 7: Fluorescence Lifetimes
http://www.iss.com/resources/reference/data_tables/LifetimeDataFluorophores.html

## Slide 8: Exctinction
Using Beer’s law (Beer-Lambert law) for light travelling through a curvette thickness d cm containing n molecules/cm3
  ln (Io/I) = nd
    where Io and I are the light entering and leaving and  is the molecular property called the absorption cross section
Now we can state that
  ln (Io/I) = nd where C is the concentration and  is the absorption coefficient which reflects the capacity of the absorbing substance to absorb light
If there are n (molecules/cm3 ; d in cm,  must be in cm2 so if  is in cm2/mol, C must be in mol/cm3 so C= /103
giving
  log10 (Io/I) = d = A
  where A is the absorbance or optical density
  and  is the decadic molar exctinction coeficient in dm3mol-1cm-1
Shapiro p 86

## Slide 9: Absorbance
O.D. units or absorbance is expressed in logarithmic terms so they are additive.
e.g. an object of O.D. of 1.0 absorbs 90% of the light. Another object of O.D. 1.0 placed in the path of the 10% of the light 10% of this light or 1% of the original light is transmitted by the second object
It is possible to express the absorbance of a mixture of substances at a particular wavelength as the sum of the absorbances of the components
You can calculate the cross sectional area of a molecule to determine how efficient it will absorb photons. The extinction coefficient indicates this value
3rd ed. Shapiro p 87
4th Ed. Shapiro p 111
Notes: OD – or optical density – or absorbance is important as it impacts many aspects of cytometyr. One is that we do need to appreciate how we reduce the amount of light when we need to. One way is to use optical filters that reduce intensity in anon-spectrally dependent fashion. That means, they reduce all bands pretty much at the same rate. Absorbance is expressed in logarithmic terms so its easy to undestand how adding neutral density filters works – so for an OD of 1.0, 90% of the light is absorbed….and if we add another filter exactly the same, we again reduce this by a factor of 10, so we end up with 1% and so on.
Wehen we look at how absorbance works in molecules, we can caluclate the cross sectional area of the molecule and this will tell us how efficent that mlecule will absorb photons thrown at it from the laser excitation. This value is the extinction coefficent.

## Slide 10: Parameters
Extinction Coefficient
   refers to a single wavelength (usually the absorption maximum) …the cross sectional area of a molecule determines how efficient it will absorb photons
Quantum Yield
  Qf is a measure of the integrated photon emission over the fluorophore spectral band
At sub-saturation excitation rates, fluorescence intensity is proportional to the product of  and Qf


## Slide 11: Fluorescence
Quantum Yield
photons emitted
photons absorbed
Q =
Fluorescence Lifetime ()
  - is the time delay between the absorbance and the emission
kr
kr + knr
=
1
kr + knr
=

R - reactant molecules)

## Slide 12: Fluorescence
Photon emission as an electron returns from an excited state to ground state
Ground State: Electrons are in positions of lowers energy
Excited State: Electron is in a temporarily position of higher energy and is very unstable and desperately wants to return to ground state in positions of lowers energy
Stable
Unstable

## Slide 14: Fluorescence
Excitation Spectrum
  Intensity of emission as a function of exciting wavelength (this is the absorbance component)
Super efficient excitation here
Inefficient here
OK
efficiency here

## Slide 15: About excitation
Intensity of emission as a function of exciting wavelength
More efficient excitation results in more signal!!

## Slide 16: Fluorescence
Chromophores are components of molecules which absorb light
They are frequently aromatic rings
or similar
FITC
Calcein
Hoechst

## Slide 17: Fluorescence
The wavelength of absorption is related to the size of the chromophores
Smaller chromophores, higher energy (shorter wavelength)
Texas Red has to be excited up here
DAPI
Texas Red
Chemical formula‎: ‎C31H29ClN2O6S2
Molar mass‎: ‎625.15 g·mol−1
CAS Number‎: ‎82354-19-6
ChEBI‎: ‎CHEBI:51247
CAS Number‎: ‎28718-90-3
Molar mass‎: ‎277.33 g·mol−1
Chemical formula‎: ‎C16H15N5

## Slide 18
© Microsoft Corp, 1995
Electromagnetic Spectrum
Only a very small region within the ES
is used for flow cytometry applications
Notes: Arc lamps and lasers
for flow cytometry
restricted to this region

## Slide 19
Properties of Fluorescent Molecules
Large extinction coefficient at the region of excitation
High quantum yield
Optimal excitation wavelength
Photostability
Excited-state lifetime
Minimal perturbation by probe

## Slide 20: Simplified Jablonski Diagram
S0
S’1
Energy
S1
hvex
hvem

## Slide 21: Fluorescence
ENERGY
S0
S1
S2
T2
T1
ABS
FL
I.C.
ABS - Absorbance S 0.1.2 - Singlet Electronic Energy Levels
FL - Fluorescence T 1,2 - Corresponding Triplet States
I.C.- Nonradiative Internal Conversion IsC - Intersystem Crossing PH - Phosphorescence
IsC
IsC
PH
[Vibrational sublevels]
Jablonski Diagram
Vibrational energy levels
Rotational energy levels
Electronic energy levels
Singlet States
Triplet States
3rd Ed. Shapiro p 87
4th Ed. Shapiro p 112

## Slide 22: Fluorescence
Stokes Shift
  is the energy difference between the lowest energy peak of absorbance and the highest energy of emission
495 nm
518 nm
Stokes Shift is 20 + nm
Fluorescein
molecule
Fluorescnece Intensity
Wavelength

## Slide 23: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  eg. UV light from sun - this causes the sunburn, not the red visible light

## Slide 24: Fluorescence Excitation Spectra
Intensity
related to the probability of the event
Wavelength
  the energy of the light absorbed or emitted

## Slide 25: Conclusions
Dye molecules must be close to but below saturation levels for optimum emission
Fluorescence emission is longer than the exciting wavelength
The energy of the light increases with reduction of wavelength

## Slide 26: Relative absorbance of phycobiliproteins
Protein | 488nm % ABS | 568nm % ABS | 633nm % ABS
B-phycoerytherin | 33 | 97 | 0
R-phycoerytherin | 63 | 92 | 0
allophycocyanin | 0.5 | 20 | 56
Data from Molecular Probes Website
Phycobiliproteins are stable and highly soluble proteins derived from cyanobacteria and eukaryotic algae with quantum yields up to 0.98 and molar extinction coefficients of up to 2.4 × 106 (ABS=Absorbance)

## Slide 27: Allophycocyanin (APC)
Protein
632.5 nm (HeNe)
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm

## Slide 28: Some Common Laser lines
Ethidium
PE
cis-Parinaric acid
Texas Red
PE-TR Conj.
PI
FITC
600 nm
300 nm
500 nm
700 nm
400 nm
457
350
514
610
632
488
Common Laser Lines
325nm

## Slide 29: Quenching, Bleaching & Saturation
Quenching is when excited molecules relax to ground states via nonradiative pathways avoiding fluorescence emission (vibration, collision, intersystem crossing)
Molecular oxygen quenches by increasing the probability of intersystem crossing
Polar solvents such as water generally quench fluorescence by orienting around the exited state dipoles
3rd Ed. Shapiro p 90
4th Ed. Shapiro p 115

## Slide 30: Photobleaching
Defined as the irreversible destruction of an excited fluorophore
Photobleaching is not a big problem for flow cytometry as long as the time window for excitation is very short (a few hundred microsconds)

## Slide 31: Fluorescence

## Slide 32: Photobleaching example from microscopy
FITC - at 4.4 x 1023 photons cm-2 sec-1 FITC bleaches with a quantum efficiency Qb of 3 x 10-5
Therefore FITC would be bleaching with a rate constant of 4.2 x 103 sec-1 so 37% of the molecules would remain after 240 sec of irradiation.
In a single plane, 16 scans would cause 6-50% bleaching

## Slide 33: Excitation Saturation
The rate of emission is dependent upon the time the molecule remains within the excitation state (the excited state lifetime f)
Optical saturation occurs when the rate of excitation exceeds the reciprocal of f
(Microscopy example) In a scanned image of 512 x 768 pixels (400,000 pixels) if scanned in 1 second requires a dwell time per pixel of 2 x 10-6 sec.
Molecules that remain in the excitation beam for extended periods have higher probability of interstate crossings and thus phosphorescence – fortunately this rarely happens in flow cytometry
Usually, increasing dye concentration can be the most effective means of increasing signal when energy is not the limiting factor (i.e. laser based confocal systems)

## Slide 34: Phosphorescence
Following absorption, molecules can relax via a non-radiative transition to the T1 rather than the S1 state - this is called an intersystem crossing
While it is forbidden it does happen and has a low probability and takes a longer time - the energy dissipated is called phosphorescence
Phosphorescence has a longer lifetime than fluorescence (milliseconds rather than femtoseconds)
Phosphorescence generally occurs at longer wavelengths than fluorescence because the energy difference between S0 and T1 is lower
3rd Ed Shapiro p 88
4th Ed. Shapiro p 113

## Slide 35: Fluorochrome excitation and emission spectra
Typical fluorchromes
FITC
PE
PerCP-Cy5.5
PE-Cy7
http://m.bdbiosciences.com/us/s/spectrumviewer

## Slide 36: Excitation - Emission Peaks
Fluorophore EXpeak EM peak
% Max Excitation at
488 568 647 nm
FITC 496 518 87 0 0
Bodipy 503 511 58 1 1
Tetra-M-Rho 554 576 10 61 0
L-Rhodamine 572 590 5 92 0
Texas Red 592 610 3 45 1
CY5 649 666 1 11 98
Note: You will not be able to see CY5 fluorescence under the regular fluorescent microscope because the wavelength is too long.
You should select fluorophores that match the laser excitation you have – and try to make sure they are not excited by other lines

## Slide 37: Excitation Efficiency

## Slide 38: The problem of spectral overlap
From BD Spectra viewer

## Slide 39
From BD Spectra viewer

## Slide 43: Resonance Energy Transfer
Resonance energy transfer can occur when the donor and acceptor molecules are less than 100 Å of one another (preferable 20-50 Å)
Energy transfer is non-radiative which means the donor is not emitting a photon which is absorbed by the acceptor
Förster (Fluorescence) RET (FRET) can be used to spectrally shift the fluorescence emission of a molecular combination.
3rd Ed. Shapiro p 90
4th Ed. Shapiro p 115

## Slide 44: FRET properties
Isolated donor
Donor distance too great
Donor distance correct

## Slide 45: Fluorescence Förster Resonance Energy Transfer (FRET)
Intensity
Wavelength
Absorbance
of donor
DONOR
Absorbance
of acceptor
molecule
Fluorescence
of donor
Fluorescence
of acceptor after
excitation by donor
ACCEPTOR
Molecule 1
Molecule 2
Laser line 488nm
550nm
600nm
650nm

## Slide 46: Note added of interest
First synthesis of fluorescein:
A. Von Baeyer, Ber Deutch Chem Ges 4:555. 1871
Recent uses of fluorescent protein molecules FUCCI
Fluorescent Ubiqutination based cell cycle Indicator
(fluorescent protein-based system that employs both a red (RFP) and a green (GFP) fluorescent protein fused to different regulators of the cell cycle: cdt1 and geminin. )
From Invitrogen website:
Developed by Miyawaki and colleagues1, the fluorescence ubiquitination cell cycle indicator (FUCCI), is a genetically encoded, two-color (red and green), indicator that lets you follow cell division within a cell population. We have incorporated the FUCCI genetic constructs into the powerful BacMam gene delivery system, creating a simple and efficient method for labeling cells and following their division:
Population-level, real-time cell cycle phase determination
Easy and efficient labeling through BacMam gene delivery system
http://cell.firstlightera.com/EN/Microsites/1/Invitrogen/PremoFUCCI
1. Sakaue-Sawano et al. (2008) Cell 132:487–498
WIKI
Johann von Baeyer
1835-1917
1905 Nobel Prize in Chemistry
Fluorescein is an organic dye also known as D & C Yellow No. 7 and many other names. A. von Baeyer first synthesized it in 1871 from phthalic anhydride and resorcinol and called it resorcinphthalein. It is not soluble in water or many organic solvents, but its disodium salt (fluorescein sodium, uranine, D & C Yellow No. 8) is highly water-soluble.
https://www.acs.org/content/acs/en/molecule-of-the-week/archive/f/fluorescein.html
(chemical formula C20H12O5)
Lifetime 3 ns MWt 332.32
Fluorescein isothiocyanate

## Slide 47: Raman Scatter
A molecule may undergo a vibrational transition (not an electronic shift) at exactly the same time as scattering occurs
This results in a photon emission of a photon differing in energy from the energy of the incident photon by the amount of the above energy - this is Raman scattering.
The dominant effect in flow cytometry is the stretch of the O-H bonds of water. At 488 nm excitation this would give emission at 575-595 nm
Can use surface enhance Raman Scattering
3rd Ed. Shapiro p 93
4th Ed. Shapiro p 118
Can use antibody-targeted, PEG-coated surface-enhanced Raman scattering (SERS) Au nanoparticles are used to label antigens
Langmuir 2013, 29, 6, 1908-1919

## Slide 49: Probes for Proteins
FITC 488 525
PE 488 575
APC 630 650
PerCP™ 488 680
Cascade Blue 360 450
Coumerin-phalloidin 350 450
Texas Red™ 610 630
Tetramethylrhodamine-amines 550 575
CY3 (indotrimethinecyanines) 540 575
CY5 (indopentamethinecyanines) 640 670
Probe Excitation Emission

## Slide 50: Specific Organelle Probes
BODIPY Golgi 505 511
NBD Golgi 488 525
DPH Lipid 350 420
TMA-DPH Lipid 350 420
Rhodamine 123 Mitochondria 488 525
DiO Lipid 488 500
diI-Cn-(5) Lipid 550 565
diO-Cn-(3) Lipid 488 500
Probe Site Excitation Emission
BODIPY - borate-dipyrromethene complexes
NBD - nitrobenzoxadiazole
DPH - diphenylhexatriene
TMA - trimethylammonium

## Slide 51: Fluorescence Overlap

## Slide 52: Fluorescence Overlap

## Slide 53: Fluorescence Overlap
a
b
Overlap of FITC fluorescence in PE PMT
Overlap of PE fluorescence in FITC PMT
This is your bandpass filter

## Slide 54: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  eg. UV light from sun - this causes the sunburn, not the red visible light
The spectrum is independent of precise excitation line but the intensity of emission is not

## Slide 55: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 56: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 57: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 58
J. Paul Robinson, Class lecture notes, BMS 631
Excitation of 3 Dyes with emission spectra

## Slide 59: Change of Excitation
J. Paul Robinson, Class lecture notes, BMS 631

## Slide 60: Tandem conjugates

## Slide 61: APC is exited nicely by 632 nm
From BD Spectra viewer

## Slide 62: APC and the Tandem
From BD Spectra viewer

## Slide 63: And now with 3 probes from 632 nm
From BD Spectra viewer

## Slide 64: And now with 3 probes from 561 nm

## Slide 65: Change Excitation – nothing!
From BD Spectra viewer

## Slide 66: PE-tandems
From BD Spectra viewer

## Slide 67: PE Tandems with 561 Excitation – more efficient
From BD Spectra viewer

## Slide 68: We can learn about dyes and filters by going to several sites to reveal properties
BioLegend: https://www.biolegend.com/spectraanalyzer
BD: http://www.bdbiosciences.com/us/s/spectrumviewer
Thermo: https://www.thermofisher.com/us/en/home/life-science/cell-analysis/labeling-chemistry/fluorescence-spectraviewer.html
Chroma: https://www.chroma.com/spectra-viewer
Sigma: https://www.sigmaaldrich.com/labware/learning-center/spectral-viewer.html
Evrogen: http://evrogen.com/spectra-viewer/viewer.shtml
Expedeon: https://www.expedeon.com/resources/applications/flow-cytometry/new-fluorescence-spectra-viewer/
Fluorophoresorg: http://www.fluorophores.tugraz.at/substance/
Semrock: https://searchlight.semrock.com

## Slide 69: DNA analysis – using PI
Intensity
Cell #
Other basic fluorescence applications- PI is an intercalating agent
The basis for this working is that there is a relationship between the number of molecules of PI and the binding to DNA is stoichiometric– thus we can generally relate dye concentration to the amount of DNA

## Slide 70: Quenching, Bleaching & Saturation
Quenching is when excited molecules relax to ground states via nonradiative pathways avoiding fluorescence emission (vibration, collision, intersystem crossing)
Molecular oxygen quenches by increasing the probability of intersystem crossing
Polar solvents such as water generally quench fluorescence by orienting around the exited state dipoles
Shapiro p 90

## Slide 71: Lecture Summary
Light and Matter
Absorption
Fluorescence
From this lecture you should understand:
  The nature of fluorescence molecules
  How fluorescence is generated
  Why molecules have different excitation and emission
  What Resonance Energy Transfer is
  What quantum yield is
  Some common fluorescent molecules
  What fluorescence overlap is
  How excitation wavelength impacts emission