## Slide 1: BMS 632 – Lecture 6 – Optical Filters & light manipulation
J. Paul Robinson SVM Professor of Cytomics Professor of Biomedical Engineering Purdue University
Reading materials: 4th Ed. Shapiro pp 119-125)
Lynn Hall
Lab: Room G221
Office G140

## Slide 2: Lecture Goals & Learning Objectives
This lecture is intended to describe the nature and function of optical systems
It will describe how optical filters are made and operate
What the properties of optical filters are
When filters should be used
What problems and issues must be taken into consideration

## Slide 3: Optics - Filter Properties
Using filters with laser systems demands higher quality and more stringent management because of the high intensity of the laser which can be 6 orders of magnitude higher than a fluorescence signal
Therefore filters must have strict limits for blocking and allowing light to pass
Rejection of light is crucial because low levels of light from lasers will override your low level signals and this rejection level is established by a % transmission of light

## Slide 4: Optics - Filter Properties
Long pass filters transmit wavelengths above a cut-on wavelength
Short pass filters transmit wavelengths below a cut-off wavelength
Band pass filters transmit wavelengths in a narrow range around a specified wavelength
  Band width can be specified
Neutral Density filter is a nondiscriminant intensity reducing filter
Absorption Filter is colored glass that absorbs unwanted light
Dichroic Mirror/Filter is a filter that passes some bands while reflecting other bands

## Slide 5: Nomenclature and Conventions
Excitation filter (D480/30x) -- Center wavelength is at 480nm; full bandwidth is 30 [ = +/- 15]. In some cases for which the band width is not specified the letter "x" is used to define the filter as an excitation filter. This is generally used for narrow band UV excitation filters, i.e. d340x.
Dichroic beamsplitter (505DCLP) -- Cut-on wavelength is 505nm for this dichroic longpass filter.
Emission filter (D535/40m) -- Center wavelength here is at 535nm; full bandwidth is 40nm [ = +/- 20].
LP -- indicates a longpass filter which transmits wavelengths longer than the cut-on and blocks shorter wavelengths
SP -- indicates a shortpass filter which transmits wavelengths shorter than the cut-on, and blocks longer wavelengths
DCLP -- dichroic longpass
DCXR -- dichroic long pass, extended reflection
DCXRU -- dichroic longpass, extended reflection including the UV
PC -- polychroic beamsplitter. Beamsplitter that reflects and transmits more than two bands of light.
GG -- Green Glass. Longpass absorption glass from Schott Glassworks with cut-on wavelengths in the violet and blue-green regions.
OG -- Orange Glass. Longpass absorption glass from Schott Glassworks with cut-on wavelengths in the green, yellow and orange regions.
RG -- Red Glass. Longpass absorption glass from Schott Glassworks with cut-on wavelengths in the red and far red regions.
x -- excitation filter; bs -- beamsplitter ; m -- emission filter

## Slide 6: Optics – Dichroic Mirrors
Dichroic mirrors allow some light to pass (transmitted light) and bounce some light (reflected light)
Mostly we use Dichroics at 45 degrees because that is the most effective design
(google dictionary)

## Slide 7: Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus a sharp objet casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution
3rd Ed. Shapiro p 83

## Slide 8: Interference in Thin Films
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light
3rd Ed. Shapiro p 82

## Slide 9: Optical filters
Interference filters: Dichroic, Dielectric, reflective filters…….reflect the unwanted wavelengths
Absorptive filters: Colored glass filters…..absorb the unwanted wavelengths

## Slide 10: Interference filters
They are composed of transparent glass or quartz substrate on which multiple thin layers of dielectric material, sometimes separated by spacer layers
Permit great selectivity

## Slide 11: Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light

## Slide 12
Standard Band Pass Filters
Transmitted Light
White Light Source
530 nm/20 BandPass Filter
520 -540 nm Light
Example: FITC excited by a 488nm laser and a 530/20 filter
This is the “band” of light

## Slide 13: A 575 nm band pass (10 nm & 10 nm wide band)

## Slide 14: A 575 nm band pass (10 nm wide band)
10 nm
center
FWHM
This means that the 575 nm is in the center of the band, and that 5 nm on either side are going to be transmitted – the width of this filter is effectively 20 nm and so it will be a filter that transmits from 570 nm to 580 nm

## Slide 15: Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters

## Slide 16: Long Pass filter
Transmission Curve

## Slide 17: Dichroics
They used to direct light in different spectral region to different detectors.
They are interference filters , long pass or short pass.
"dichroic" Di- is Greek for two, and -chroic is Greek for color - from Greek dikhroos, bicolored

## Slide 18: Optical Filters
Dichroic Filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source

## Slide 19: Dichroic Filters
Transmitted
Light
Reflected
Light
Filter acting as a DICHROIC

## Slide 20: Construction of Filters
Filter
components
Single Optical
filter
“optical
glue” or
mostly filters
are spatter
Coated in a
vacuum
Interference filters

## Slide 21: Transmission determination
Constructive and destructive interference occurs between reflections from various layers
Transmission determined by :
  thickness of the dielectric layers
  number of these layers
  angle of incident light on the filters

## Slide 22: Absorptive filters
Such as colored glass filters which absorb unwanted light.
Consist of dye molecules uniformly suspended in glass or plastic.
Remove much more of the unwanted light than do the interference filters
Will often fluoresce (not good!)

## Slide 23: Definitions
400 nm
500 nm
600 nm
700 nm
Wavelength
Half Max
Max
Full Width at
Half Max
(FWHM)
% Transmission
Wavelength
0
100

## Slide 24: Filters transmission
Bandpass filters: characterized by there
T max and (the Full Width at Half Maximum) FWHM
Notch filters are band pass filters in the upside down position
Long pass and Short pass filters: characterized by their T max and cut-on, cut-off wavelength

## Slide 25: Fluorescein - (FITC)
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Band Pass Filter

## Slide 26: Fluorescein - (FITC)
400 nm
500 nm
600 nm
700 nm
Wavelength
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Band Pass Filter

## Slide 27: Fluorescein-Phycoerytherin - (FITC-PE)
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Band Pass Filters

## Slide 28: Using a Band pass filter correctly
https://www.omegafilters.com/curvo2/index.php

## Slide 29
Source: https://www.omegafilters.com/

## Slide 30: Finding the right filter -
Source: http://www.chroma.com/index.php?option=com_products&Itemid=53#
Exciter

## Slide 31
Source: from Chroma website
Exciter

## Slide 33
Typical Emission scan
500

## Slide 34: Laser Blocking Filters
525 nm LEDs

## Slide 35: Interference filters advantages
They can be used as reflectors in two and three color analysis.
They usually do not themselves produce fluorescence.
They are available in short pass versions.
They are excellent as primary barrier filters.

## Slide 36: The output of a band pass filter

## Slide 39: Filter 525-10
Just because the filter says “525-10” does not
Mean that is its only transmission!!
What is all this?

## Slide 41
CFP-YFP filter ex - em
CFP-YFP-Hc-Red filter ex - em

## Slide 42: Filter damage
Moisture
Humidity can cause filters to become “cloudy” or for layers to be physically damaged – this will change the property of the filter.
Laser damage – if a filter is an excitation filter, it can be destroyed by the very laser it is designed to work with – so check excitation filters if they are used (often in imaging)
Old age!! – Who knows what happens to filters that are “just old” – they should be tested regularly (at least once every couple of years.

## Slide 44: 395 nm bad filter

## Slide 45: 408 nm Long Pass

## Slide 46: Optical Filters Worksheet
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light
Standard Band Pass Filters (BP)
Standard Long Pass Filters (LP)
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters (SP)
Dichroic filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source
Standard Dichroic Filters (DC)
Create the following filters
Show filters for these dyes
FITC
APC
DAPI
Cy5
PE
FITC
%T
λ
%T
λ
%T
λ
%T
λ
515 BP
575 DC/LP
680 DC/LP
620 SP
%T
λ
520nm

## Slide 47
http://evrogen.com/spectra-viewer/viewer.shtml
Example

## Slide 48: Chroma Spectra Viewer
https://www.chroma.com/spectra-viewer?fluorochromes=10484&sets=13593
Example

## Slide 49
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/labeling-chemistry/fluorescence-spectraviewer.html
Example
Mercury Lamp Excitation
633nm Laser Excitation

## Slide 50: BD Spectrum Viewer
http://www.bdbiosciences.com/us/s/spectrumviewer
Example

## Slide 51
Dyes
Lasers
Example

## Slide 52: We can learn about dyes and filters by going to several sites to reveal properties
BioLegend: https://www.biolegend.com/spectraanalyzer
BD: http://www.bdbiosciences.com/us/s/spectrumviewer
Thermo: https://www.thermofisher.com/us/en/home/life-science/cell-analysis/labeling-chemistry/fluorescence-spectraviewer.html
Chroma: https://www.chroma.com/spectra-viewer
Sigma: https://www.sigmaaldrich.com/labware/learning-center/spectral-viewer.html
Evrogen: http://evrogen.com/spectra-viewer/viewer.shtml
Expedeon: https://www.expedeon.com/resources/applications/flow-cytometry/new-fluorescence-spectra-viewer/
Fluorophoresorg: http://www.fluorophores.tugraz.at/substance/
Semrock: https://searchlight.semrock.com

## Slide 53: Interference filters: disadvantages
Lower blocking properties
Reduced passing properties
Their reflecting and passing properties are not absolute, this should be considered while dealing with multiple wavelengths

## Slide 54: Conclusions
Optical filters play a critical role in selecting appropriate bands of fluorescence emission
`Each filter will be responsible for some loss in transmission
Absorbance filters may actually cause fluorescence
Filter efficiency may be important in complex optical systems