## Slide 1: BMS 631 - LECTURE 3 Light and Matter
Lynn Hall, G221
Purdue University
Office: 494 0757
email; wombat@purdue.edu
WEB http://www.cyto.purdue.edu
4th Ed Shapiro p 100-130
J.Paul Robinson Distinguished Professor of Cytometry College of Veterinary Medicine & Professor of Biomedical Engineering College of Engineering Purdue University

## Slide 2: Light and Matter
Energy
  joules, radiant flux (energy/unit time)
  watts (1 watt=1 joule/second)
Angles
  steradians - sphere radius r - circumference is 2r2; the angle that intercepts an arc r along the circumference is defined as 1 radian. (57.3 degrees) a sphere of radius r has a surface area of 4r2. One steradian is defined as the solid angle which intercepts as area equal; to r2 on the sphere surface
Shapiro p 75

## Slide 3: Terms
Side scatter, forward angle scatter, cell volume, coulter volume:
Understand light scattering concepts; intrinsic and extrinsic parameters
Photometry:
Light - what is it - wavelengths we can see 400-750 nm, most sensitive around 550 nm. Below 400 nm essentially measuring radiant energy. Joules (energy) radiant flux (energy per unit time) is measured in watts (1 watt=1 joule/second).
Steradian (sphere radius r has surface area of 4 pr2; one steradian is defined as that solid angle which intercepts an area equal to r2 on the surface.
Mole - contains Avogadro's number of molecules (6.02 x 1023) and contains a mass in grams = molecular weight. Photons - light particles - waves - Photons are particles which have no rest mass - pure electromagnetic energy - these are absorbed and emitted by atoms and molecules as they gain or release energy. This process is quantized, is a discrete process involving photons of the same energy for a given molecule or atom. The sum-total of this energy gain or loss is electromagnetic radiation propagating at the speed of light (3 x 108 m/s). The energy (joules) of a photon is
E=hn and E=hn/l [n-frequency, l-wavelength, h-Planck's constant 6.63 x 10-34 joule-seconds]
Energy - higher at short wavelengths - lower at longer wavelengths.

## Slide 4: Photons and Quantum Theory
Photons
  particles have no rest mass - composed of pure electromagnetic energy - the absorption and emission of photons by atoms and molecules is the only mechanism for atoms and molecules can gain or lose energy
Quantum mechanics
  absorption and emission are quantized - i.e. discrete process of gaining or losing energy in strict units of energy - i.e. photons of the same energy (multiple units are referred to as electromagnetic radiation)
Energy of a photon
  can be computed from its frequency ()
  in hertz (Hz) or its wavelength (l) in meters from
        E=h and E=hc/
 = wavelength
h = Planck’s constant
(6.63 x 10-34 joule-seconds
c = speed of light (3x108 m/s)
3rd ed Shapiro p 76

## Slide 5: Light – Particles and waves - Reflection
Diagrams from: http://micro.magnet.fsu.edu/primer/java/particleorwave/reflection/index.html

## Slide 6: Light – Particles and waves -
Refraction
Diffraction
Diagrams from: http://micro.magnet.fsu.edu/primer/java/particleorwave/reflection/index.html

## Slide 7: Laser power
One photon from a 488 nm argon laser has an energy of
  E= 6.63x10-34 joule-seconds x 3x108
To get 1 joule out of a 488 nm laser you need 2.45 x 1018 photons
1 watt (W) = 1 joule/second so a 10 mW laser at 488 nm is putting out 2.45x1016 photons/sec
E=h and E=hc/
= 4.08x10-19 J
Shapiro p 77
(lasers will be discussed in more detail in a later lecture)

## Slide 8: What about a HeCd UV laser? (325 nm)
  E= 6.63x10-34 joule-seconds x 3x108
= 6.12 x 10-19 J so 1 Joule at 325 nm = 1.63x1018 photons
What about a He-Ne laser? (633 nm)
  E= 6.63x10-34 joule-seconds x 3x108
= 3.14 x 10-19 J so 1 Joule at 633 nm = 3.18x1018 photons
Shapiro p 77
A photon has energy
A photon has energy
325 x 10-3
633 x 10-3

## Slide 9: Polarization and Phase: Interference
Electric and magnetic fields are vectors - i.e. they have both magnitude and direction
The inverse of the period (wavelength) is the frequency in Hz
Shapiro p 78
Wavelength (period T)
Axis of
Magnetic Field
Axis of Propagation
Axis of Electric Field

## Slide 10: Interference – Critical to understanding optical filters
Constructive
Interference
Destructive
Interference
A
B
C
D
A+B
C+D
Amplitude
0o
90o
180o
270o
360o
Wavelength
Figure modified from Shapiro “Practical Flow Cytometry” Wiley-Liss, p79
Here we have a phase difference of 180o (2 radians) so the waves cancel each other out
The frequency does not change, but the amplitude is doubled

## Slide 11: Interference in Thin Films
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light
Shapiro p 82

## Slide 12: Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus, a sharp object casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution
Shapiro p 83
http://micro.magnet.fsu.edu/primer/java/interference/doubleslit/index.html
Thomas Young’s double split experiment in 1801

## Slide 13: Light Scatter
Materials scatter light at wavelengths at which they do not absorb
If we consider the visible spectrum to be 350-750 nm then small particles (< 1/10 ) scatter rather than absorb light
For small particles (molecular up to sub micron) the Rayleigh scatter intensity at 0o and 180o are about the same
Shapiro p 79
For larger particles (i.e. size from 1/4 to tens of wavelengths) larger amounts of scatter occur in the forward not the side scatter direction - this is called Mie Scatter (after Gustav Mie) - this is how we basically come up with the concept that forward scatter is somewhat related to size – we will also talk later about Axial Light Loss (ALL) which is considered to be a more accurate determination of particle size
Gustav Mie 1969-1957
From Wikipedia

## Slide 14: Light Scatter video….. https://www.youtube.com/watch?v=IOOtw3TQXE8
Video 6:09 Mins

## Slide 15: The “laser” volume vs the “particle” volume
10 μ
The fundamental problem is that microparticles are detected within the noise region of virtually all flow cytometers; despite a number of attempts to measure them via FC, with some level of success, all have identified significant problems because of the failure of cytometers to adequately separate microparticles from noise. It is almost absurd to consider a minimum of a 12-order differentiation acceptable. There are increasing numbers of papers making unsustainable claims about the effectiveness of FC for sub micron–size particles, most using fluorescence-based triggering that cannot guarantee anything about the size or even the number of particles within the excitation volume
Flat top laser
- this tries to make
the illumination
as even as possible
Larger elliptical laser
line trying to increase
the size of the illumination
volume and flattening it out
Small elliptical laser
line trying to decrease
the size of the illumination
volume to find small
particles

## Slide 16: Light Scatter and Particle size
Dia (ratio 1) | Volume 4/3 πr3 | Ratio | Surface Area | Ratio particle to laser vol
10 microns =1 | 4.1887902E-6 μ3 | 1 | 1.25663706E-3 μ2 | 1
1 micron 1/10 | 4.1887902E-9 μ3 | 1E+3 | 1.25663706E-5 μ2 | 1E+2
500 nm 1/50 | 5.23598776E-10 μ3 | 8E+3 | 3.14159265E-6 μ2 | 4E+2
100 nm 1/100 | 4.1887902E-12 μ3 | 1E+6 | 1.25663706E-7 μ2 | 1E+4
50 nm 1/500 | 5.23598776E-13 μ3 | 8E+6 | 3.14159265E-8 μ2 | 4E+4
10 nm 1/1000 | 4.1887902E-15 μ3 | 1E+9 | 1.25663706E-9 μ2 | 1E+6
Having some idea of the size of a particle, will give you knowledge of what the detection limits might be of a flow cytometer
A fundamental problem with small particles is that they are MUCH smaller than the laser illumination volume

## Slide 17: Rayleigh Scatter
Molecules and very small particles do not absorb, but scatter light in the visible region (same frequency as the excitation)
Rayleigh scattering is directly proportional to the electric dipole and inversely proportional to the 4th power of the wavelength of the incident light
The sky looks blue because the gas molecules scatter more light at shorter (blue) rather than longer wavelengths (red)

## Slide 18: Refraction & Dispersion
Light is “bent” and the resultant colors separate (dispersion).
Red is least refracted, violet most refracted.
dispersion
Short wavelengths are “bent” more than long wavelengths
ref
rac
tion

## Slide 19: Reflection and Refraction
Snell’s Law: The angle of reflection (Ør) is equal to the angle of incidence (Øi) regardless of the surface material
The angle of the transmitted beam (Øt) is dependent upon the composition of the material
Shapiro p 81
t
i
r
Incident Beam
Reflected Beam
Transmitted
(refracted)Beam
n1 sin Øi = n2 sin Øt
The velocity of light in a material of refractive index n is c/n

## Slide 20: Brewster’s Angle
Brewster’s angle is the angle at which the reflected light is linearly polarized normal to the plane incidence
At the end of the plasma tube, light can leave through a particular angle (Brewster’s angle) and essentially be highly polarized
Maximum polarization occurs when the angle between reflected and transmitted light is 90o
    thus Ør + Øt = 90o
    since sin (90-x) = cos x
    Snell’s provides (sin Øi / cos Øi ) = n2/n1
    Ør is Brewster’s angle
Shapiro p 82
Ør = tan -1 (n2/n1)
For a glass medium (n2 ≈ 1.5) in air (n1 ≈ 1), Brewster's angle for visible light is approximately 56°, while for an air-water interface (n2 ≈ 1.33), it is approximately 53°. Since the refractive index for a given medium changes depending on the wavelength of light, Brewster's angle will also vary with wavelength.
Thanks Wiki
David Brewster
Thanks Wiki
Solid state lasers used today are linearly polarized whereas the ion lasers that were once in flow cytometry were not

## Slide 21: Brewster’s Angle
Seen on the laser tubes from an argon laser
Brewster mirror
at 56° on ion lasers

## Slide 22: Fluorescence Absorption
Basic quantum mechanics requires that molecules absorb energy as quanta (photons) based upon a criteria specific for each molecular structure
Absorption of a photon raises the molecule from ground state to an excited state
Total energy is the sum of all components (electronic, vibrational, rotational, translations, spin orientation energies) (vibrational energies are quite small)
The structure of the molecule dictates the likely-hood of absorption of energy to raise the energy state to an excited one (that is why some molecules are fluorescent and other are not!
Shapiro p 84
(Discussed more in lecture on fluorescence)

## Slide 23: Fluorescence Lifetime
Absorption associated with electronic transitions (electrons changing states) occurs in about 1 femptosecond (10-15 s)
Fluorescence lifetime is defined as the time in which the initial fluorescence intensity of a fluorophore decays to 1/e (approx 37 percent) of the initial intensity
The lifetime of a molecule depends on how the molecule disposes of the extra energy
Because of the uncertainty principle, the more rapidly the energy is changing, the less precisely we can define the energy
So, long-lifetime-excited-states have narrow absorption peaks, and short-lifetime-excited-states have broad absorption peaks
Shapiro p 85
http://micro.magnet.fsu.edu/primer/techniques/fluorescence/fluorescenceintro.html

## Slide 24: Fluorescence Lifetimes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/
Most dyes used in current flow cytometry have
a lifetime of between 0.1 and 20 ns
Quantum dots can be much longer lifetime

## Slide 25: Fluorescence Lifetimes
http://www.iss.com/resources/reference/data_tables/LifetimeDataFluorophores.html
Some common flow cytometry fluors in the
3-5 ns lifetime range

## Slide 26: So, can we measure lifetimes in Flow Cytometry?
The answer is only on some instruments that are custom made
Scientific Reports | 7:40341 | DOI: 10.1038/srep40341

## Slide 27: Extinction
Using Beer’s law (Beer-Lambert law) for light travelling through a curvette thickness d cm containing n molecules/cm3
  ln (Io/I) = nd
    where Io and I are the light entering and leaving and  is the molecular property called the absorption cross section
Now we can state that
  ln (Io/I) = nd where C is the concentration and  is the absorption coefficient which reflects the capacity of the absorbing substance to absorb light
If there are n (molecules/cm3 ; d in cm,  must be in cm2 so if  is in cm2/mol, C must be in mol/cm3 so C= /103
giving
  log10 (Io/I) = d = A
  where A is the absorbance or optical density
  and  is the decadic molar exctinction coeficient in dm3mol-1cm-1
Shapiro p 86

## Slide 28: Absorbance
O.D. units or absorbance is expressed in logarithmic terms so they are additive.
e.g. an object of O.D. of 1.0 absorbs 90% of the light. Another object of O.D. 1.0 placed in the path of the 10% of the light 10% of this light or 1% of the original light is transmitted by the second object
It is possible to express the absorbance of a mixture of substances at a particular wavelength as the sum of the absorbances of the components
You can also calculate the cross-sectional area of a molecule to determine how efficiently it will absorb photons.
The extinction coefficient indicates this value (discussed more in later lecture)
Shapiro p 87

## Slide 29: Fluorescence (More details in a later lecture)
Photon emission as an electron returns from an excited state to ground state
Fluorescence detection is a fundamental component of all flow cytometry systems
Fluorescence is very sensitive and its great advantage is that the emission is usually at a different wavelength to the excitation, so background is low

## Slide 30: Fluorescence
Quantum Yield
photons emitted
photons absorbed
Q =
Fluorescence Lifetime ()
  - is the time delay between the absorbance and the emission

## Slide 31: Fluorescence
Excitation Spectrum
  Intensity of emission as a function of exciting wavelength
Chromophores are components of molecules which absorb light
They are frequently aromatic rings

## Slide 32: Parameters (or Variables)
Extinction Coefficient
   refers to a single wavelength (usually the absorption maximum) (The extinction coefficient is determined by measuring the absorbance at a reference wavelength (characteristic of the absorbing molecule) for a one molar (M) concentration (one mole per liter) of the target chemical in a cuvette having a one-centimeter path length.)
  the intrinsic lifetime of a fluorophore is inversely proportional to the extinction coefficient, molecules exhibiting a high extinction coefficient have an excited state with a short intrinsic lifetime.
Quantum Yield
  Qf is a measure of the integrated photon emission over the fluorophore spectral band
  Expressed as ratio of photons emitted to the number of photons absorbed (zero to 1 (best)
At sub-saturation excitation rates, fluorescence intensity is proportional to the product of  and Qf

## Slide 33: Raman Scatter
A molecule may undergo a vibrational transition (not an electronic shift) at exactly the same time as scattering occurs
This results in a photon emission of a photon differing in energy from the energy of the incident photon by the amount of the above energy - this is Raman scattering.
The dominant effect in flow cytometry is the stretch of the O-H bonds of water. At 488 nm excitation this would give emission at 575-595 nm
Can use surface enhance Raman Scattering
3rd Ed. Shapiro p 93
4th Ed. Shapiro p 118
Can use antibody-targeted, PEG-coated surface-enhanced Raman scattering (SERS) Au nanoparticles are used to label antigens
Langmuir 2013, 29, 6, 1908-1919

## Slide 34: Quenching, Bleaching & Saturation
Quenching is when excited molecules relax to ground states via nonradiative pathways avoiding fluorescence emission (vibration, collision, intersystem crossing)
Molecular oxygen quenches by increasing the probability of intersystem crossing
Polar solvents such as water generally quench fluorescence by orienting around the exited state dipoles
Bleaching is defined as the irreversible destruction of an excited fluorophore
Saturation is where the fluorescent molecule is excited at its maximum lifetime – beyond this can result in molecular damage
Shapiro p 90
Will revise this in the lecture on Fluorescence

## Slide 35: Samples are used in flow cytometry
Biological materials particularly cells, blood, cell cultures, other body fluids
Microbial samples – bacteria, algae, viruses
Various liquids such as beer and wine analysis
Synthetic beads – calibration applications
Solid tissue – only when disaggregated

## Slide 36: Sheath examples
Sheath used to be saline because almost all early instruments were cell sorters – saline is required in cell sorters
BUT – saline is not good for fluidics so it really does not matter what you run in the sheath since in principle, the sheath and sample do not mix
Water is preferred to keep systems clean
Various companies sell sheath in convenient containers – this is useful as the idea sheath contains no small particles
While there is no reason you cannot make your own sheath – the biggest problem encountered is “clean” sheath – small particles are hard to remove and you will need some significant effort to filter out particles

## Slide 37: So now we have the fundamentals of flow cytometry
What can we do?
Immunofluorescence: measuring receptors/markers of cells
Clin Immun Immunopath 38:161-177, 1984
Arch Pathol Lab Med (2006) 130 (5): 680–690.
https://doi.org/10.5858/2006-130-680-CACFCI

## Slide 38: Functional Flow Cytometry – Platelet function
https://doi.org/10.1006/meth.2000.1006

## Slide 39: Analysis of NK Cells
https://doi.org/10.1007/978-1-60761-362-6_23

## Slide 40: Neutrophil Function by Flow Cytometry
https://doi.org/10.1016/S0022-1759(99)00148-9
https://doi.org/10.1016/S0091-679X(08)61733-1

## Slide 41: Plant Flow Cytometry
https://doi.org/10.1093/oxfordjournals.aob.a010312
Cytometry Pt A, Volume: 103, Issue: 12, Pages: 953-966, First published: 08 October 2023, DOI: (10.1002/cyto.a.24798)
An example of multi-step gating used to isolate 2n nuclei in a pollen sample. (A) Selection of nuclei region on a FL2-H (585 nm) by FL3-H (670 nm) scatterplot, used to gate out debris. (B) Scatterplot of FL2-A versus SSC for the same sample prior to gating. (C) and (D) show the same data after gating as in (A), with a region draw around the 2C nuclei either (C) manually or (D) using an automated cluster identification tool (“Snap-to” regions, CellQuest Pro, BD Biosciences). (E) 2C nuclei data as selected by the combination of gates in (A) and (C), displayed on a height versus area plot for pulse analysis (aggregate removal). The rectangular region contains only single nuclei. (F) Histogram of ungated data with the small peak outline in white showing the events of interest, single 2C nuclei, which are only a part of the 2C peak.
https://doi.org/10.1002/cyto.a.24798

## Slide 42: Fish Flow Cytometry (the swimming kind)
https://doi.org/10.1002/cyto.10100
https://doi.org/10.1007/s11626-012-9543-7

## Slide 43: Fish-Fish Flow Cytometry
Flow-FISH (fluorescence in-situ hybridization) is a cytogenetic technique to quantify the copy number of RNA or specific repetitive elements in genomic DNA of whole cell populations via the combination of flow cytometry with cytogenetic fluorescent in situ hybridization staining protocols
Wikipedia

## Slide 44: Cell Sorting
https://doi.org/10.1007/978-1-62703-038-0_26
https://doi.org/10.1016/S0167-7012(00)00181-0
Sperm
Bacteria
Cancer Cells
https://doi.org/10.1002/cyto.a.22022

## Slide 45: Summary
Review of nature of light
Review of fundamental features of light
Review of how de define efficiency of light
Introduction to fluorescence
Review of factors that influence light (quenching, lifetime, etc.)
A few examples of applications of flow cytometry